## Slide 2: Sample Holder
Holder modeled on that used in Meth paper ( open gap in middle, light clamping on sides) but vertical
No part of this support structure directly touches the subsequent heater components
Outline represents heated volume
Top
Side
Front
Sample
Incoming Laser
Thermocouple

## Slide 3: Heater
Steel square tube stock was coated in Sodium Silicate based paste
Open top and bottom
Wrapped in NiChrome resistive heating wire, power not shown
Tube acts as heat distributor
Open front for fluorescence entrance and exit
Top
Side
Front
Sample
Incoming Laser
Thermocouple
NiChrome Wire
Sodium Silicate

## Slide 4: Insulation (top cap not shown)
Air-gapped as insulator
Fiberglass reinforced epoxy board, PCB material
Copper outer layer (not shown, not considered important)
Top
Side
Front
Sample
Incoming Laser
Thermocouple
NiChrome Wire
Sodium Silicate
FR-4 PCB

## Slide 5: Insulation with top cap shown
Metal top to decrease chimney effect
In direct contact with heat distributor
Reduce airflow
Small hole for access
DON’T put thermocouples down this
Top
Side
Front
Sample
Incoming Laser
Thermocouple
NiChrome Wire
Sodium Silicate
FR-4 PCB

## Slide 6: Thermo position differences setup
Positions 2 and 3 are in contact with PMMA
Positions 1 and 4 are in air
Position B is at the sample spot height
A and C are above and below, respectively
All are referenced to the thermocouple as drawn
Top
Side
Front
1 2 3 4
A
B
C

## Slide 7: Results
Data taken at 35 or 40 C setpoint
Almost all spots technically inside thermo accuracy (+- 2 C)
Measurable difference though
  At least 8 samples per data point
 |  | Difference |  |  |  |  |  | 
 | 1 | STD | 2 | STD | 3 | STD | 4 | STD
A | -2.07125 | 0.91229 | -0.55274 | 0.58518 |  |  | 1.73701 | 0.23556
B | -1.13167 | 0.4922 |  |  | 1.38554 | 0.4407 |  | 
C | -0.92247 | 0.62455 |  |  | -1.16125 | 0.2817 | -1.87 | 0.40486
 |  | Diff as % of mounted Thermo |  |  |  |  |  | 
A | -4.78612 | 2.01199 | -1.38517 | 1.46467 |  |  | 4.37883 | 0.5836
B | -2.84694 | 1.18845 |  |  | 3.48456 | 1.08041 |  | 
C | -2.33066 | 1.57622 |  |  | -3.26931 | 0.77702 | -5.16996 | 1.03194

## Slide 8: Difference between Thermocouple in through top and bottom